SBC2023 – Serial data and IR remote control codes
IR protocol is NEC with a 8 bit system code and 8 bit command code. We have been testing with a common remote with a System Code of 0xef, but we can use any other code or IR protocol that is required. We suggest that a System Code is chosen that is not in common use by other equipment.

Motion	System Code	Command Code
Continuous motion while IR is received plus preset run-on (note 1)
		Hexadecimal		Decimal
Head/Backrest up	0xef	0x04		4
Head/Backrest Down		0x05		5
Leg/Foot up		0x06		6
Leg/Foot down		0x07		7
Bed Up		0x08		8
Bed Down		0x09		9
Tilt/Trandellenburg up		0x0a		10
Tilt/Trandellemburg down		0x0b		11
Go to Preset positions
Preset 1	0xef	0x10		16
Preset 2		0x11		17
Preset 3		0x12		18
Preset 4		0x13		19

Programming instructions (note 2)
Learn Preset 1		0xef	0x14	20
Learn Preset 2			0x15	21
Learn Preset 3			0x16	22
Learn Preset 4			0x17	23

System Wake-up
Wakeup (note 3)		0xef	0x00	0
Stop (note 4)			0x01	1

Notes
1. Use with systems that just give a single command burst such as speech recognition. The motor will run for a preset time each time the command is received. A series of repeated commands might be needed to reach a desired position.
2. Programming instructions must be received 3 times within 2 seconds to be valid and responded to.
3. This is a safety function, and the command must sent in response to the user giving a command such as “Bed wake up”. This will bring the bed control out of standby and be available for additional commands. The controller will revert to standby 60 seconds after the last command. The controller LCD backlight will be active while the unit is ready for commands.
4. The STOP command will always stop any motion on the bed, even if it is in standby.

Wired serial data format
Codes sent by serial data are sent 9600 BPS, 8NE as a 3 byte sequence:
Byte 1 – System code
Byte 2 – Command code
Byte 3 – Checksum consisting of simple 8 bit addition of previous 2 bytes