## Sheet1
@title	TSA Scanning / Infra Red Motorised Bed Controller 2022

@requirements
@heading	Packaging
	The system shall be packaged into a number of units that can be installed, upgraded seperately.
	There shall be exactly one Control unit, the only part that the user sees.
	The Control unit appearance shall be harmonious with a domestic setting
	The Control unit shall be able to be rebadged with a label or similar
	The Control Unit shall have the long sides of the box unencumbered to allow easy mounting.
	The Control Unit shall have connectors on end end of the case only.
	There shall be at least one Bed Interface unit, installed under the bed, that controls the bed and connects to external power supply.
	The Interface unit shall be able to be upgraded to connect to different bed systems.
	There shall be up to 4 Tilt Sensor units that can be attached to the parts of the bed frame to sense tilt.

@heading	Comms & Power
	There shall be a single cable to distribute power and comms to all units.
	The comms method shall be robust with error checking to cope with possible electrical interference that has been observed on some beds.
	The comms shall be a network able to connect to at least 16 devices.
	The comms network shall be single master / multiple slave.
	The comms shall allow easy connection of devices in a daisy chain or star fashion.
	The comms connection should use network cables with a RJ45 connector to allow easy replacement.
	The power rail shall be protected against shorts and overloads by shutting down rather than blowing a fuse.
	All units shall have a small LED that indicates power, bad comms, error state by blinking pattern for easy remote faultfinding.

@heading	Control Unit Facilities - Hardware
	The Control Unit shall receive power & data over the standard connector.
	The Control Unit shall have a 16 x 2 LCD with intensity controlled backlight.
	The Control Unit shall have provision for up to 4 touch buttons.
	The Control Unit shall have a 5” diagonal LCD with intensity controlled backlight.	Future
	The Control Unit LCD shall have up to 4 touch buttons.	Future
	Able to connect to a WiFi and connect with http to an arbitrary device.	Future
	Able to connect to a Smartphone over Bluetooth.	Future
	The Control Unit shall have a single switch input for use with a disability adapted switch.
	The Control Unit shall an IR receiver for receiving bed move commands.
	The unit shall have a short range IR transmitter to teach learning IR remotes.	Future		All units shall actively check for faults such as sensor failure, loss of comms, microcontroller lockup, memory corruption and either recover or turn off all motors.
	A RS232 “customer” serial port for accepting bed move commands.
	An externally accessible USB port for updating the software on the controller.

@heading	Control Unit Facilities – Software
@heading2	USB Port
	The firmware on the Control Unit shall be updatable via an USB port when connected to a PC.
	The Control Unit USB port shall be able to accept Bed Commands.
	The Control Unit USB port shall be able to configure the unit via a simple text based protocol.
	The Control Unit USB port shall be able to accept Bed Commands.
	The display shall:
	Show a idle message when it is not executing a command.
	Show the currently executing command.
	Display error messages until they are explicitly cleared with a button press.
	Configuration shall:
	Consist of a number of items that can take a range of numeric values either numbers or options.
	A configuration item to disable specific motors.
	A configuration item to set the number of Sensors.
	Be stored in non-volatile memory able to recover from a single byte error and detect more serious corruption.
	Set by a simple menu system with buttons to select a menu item and then to adjust the value.
	Set by the “customer” serial port using a simple protocol.
	Set by the USB serial port using a simple protocol.
	Commands shall:
	Be sourced from the IR receiver using a specific code mapping to a command.
	Be sourced from the “customer” serial port using a simple protocol.
	Be sourced from the USB serial port using a simple protocol.
	From a scanning interface using a switch plugged into the switch input.
	Have commands to drive each motor in either direction for a configurable duration.
	The duration might be different for different command sources.
	Have a command to stop all motors.
	Have a command to record position from a number of Sensor units as a position preset.
	Have a command to control the bed to return to a preset position.
	Any command can fail, but this failure is not an error, it is displayed as a message on the LCD, and might be logged, but the unit returns to a state where it is ready for the next command.
	Preset position
	For recording position, the data from the sensors shall be checked for reasonableness as much as possible.
	When moving to a preset position, there shall be options to move the motors in a specific order, or all at once.
	During moving to a preset position, the sensors shall be checked that they actually are changing value, and in the correct direction.
	During moving to a preset position, there shall be a configurable timeout.
	Set limits for those axes with Sensors installed as limits that the MBC will not drive the motor beyond.
	Logging
	There shall be a logging facility that can be configured to log all commands, or just those that result in errors.
	Size is not an issue, enough for say 1000 entries just to aid faultfinding in case of errors at a remote site.
	The log can be accessed via the “customer” serial port or the USB serial port.
	Log entries shall be tagged with a timestamp.
	What memory device is used for logging is an open issue, possibilities include a bare (no FS) SPI Flash memory device, SD card, or processor Flash memory.
	Bed Interface Unit – Hardware
	The unit shall accept a DC power input of 10V to 26V DC at a maximum of 1A.
	Nominal 12V input on a 2.5mm power inlet socket.
	Nominal 24V input on internal screw terminal with grommet in case for cable.
	DC power to be reverse input protected.
	It will distribute this as unregulated to the other units in the system.
	It will have a daughter board containing 8 relays and a selection of connectors for the most common bed controllers.
	It should be possible to add daughter boards with other interfaces such as RS485 or serial or more complex relays to suit uncommon bed controllers.
	Bed Interface Unit – Software
	Sensor Unit – Hardware
	The Sensor is required to attach to a bed frame, so ideally it should be as small as possible about twice as long as wide to give options for attachment such as two way tape of cable ties.
	It should allow daisy chaining the comms network for ease of wiring, so 2 connectors.
	Sensor Unit – Software
	The unit must check for correct operation of the actual tilt sensor device (accelerometer) as far as possible.
	The nominal range of sensing is 0° to 90°. However, the Sensor may be installed either way up, and might not be exactly horizontal, so the actual range is ±120° around horizontal.
	It is not necessary that the tilt angle be precisely calibrated, only that it can repeatably measure within the required range and that the indicated angle is monotonic against actual angle.

	Specification
	It was decided to develop a reduced capability initial version of the product to satisfy one customer who required only control from their automation system.
	So the following requirements will not be included:
	Large LCD display,
	Disability adapted switch input for scanning interface.
	IR receiver / transmitter hardware.
	WiFi & Bluetooth connectivity.

	However other components of the system will still be usable into the future.
	The customer also only required 2 sensors, on the head & knee positions.
	Packaging
	The system is to be split up physically into Control Unit, Bed Interface unit, and 2 Tilt Sensor units.
	Comms and Power
	Decided to use RS485 transport as we have experience with this as well as existing products that use it. It is also very robust electrically, and resistant to noise.
	The RS485 transmitter/receiver IC will be type MAX485E which allows up to 32 nodes on the bus, in excess of the required 16. It also is inexpensive as it is the older technology running from 5V and we have excess stock.
	However RS485 does require line level biasing in one location on the bus,
	RS485 also requires termination at the endpoints, though for moderate line lengths and speeds this is not necessary.
	Use RJ45 as the power and comms connector with two commoned sockets on each unit, allows daisy chaining units as required. The pinout is non-standard but uses one pair for the RS485 data, and two other pairs commoned for ground and unregulated DC.
	Protocol has been chosen to be MODBUS, as this is a standard master/slave protocol that we have experience with, and has adequate error checking and confirms that an addressed unit has received data.
	Error LED
	All units will have single red LED is included next to the comms sockets intended for faultfinding only.
	Normal operation is rapid blinks of around 10ms duration at around 10Hz.
	If off then there is a problem with the power rail.
	A loss of comms is indicated by a slow flash 500ms on and 500ms off.
	Any other serious error is indicated by a flash 250ms on and 2000ms off.

	Faults
	The firmware in each unit will check for the following faults:
	Enable the microcontroller watchdog to guard against lockups by requiring both the main line and an interrupt routine to both run periodically to clear the watchdog, else it resets the processor. In this case, the processor will perform a normal restart.
	Any configuration data in EEPROM will be stored as two copies each with a CRC or similar. At startup, if one is found to be corrupt, it is overwritten by the valid copy. If both are corrupt, then they are not overwritten but the device will not run and will signal an error.

	Control Unit Hardware
	A key decision in components & features is reducing development effort. So we have chosen parts that are already stock items, and are already drawn in the PCB & mechanical CAD system.
	For easy development we have selected the Arduino Mega2560 processor, as it has multiple serial ports. Also the existing software libraries for the Sensor & Bed Control units will port over directly.
	Case hardware is the same clam-shell case as used for the existing MBC. A key advantage is the removable end plates, which are easy to machine when removed. We also have jigs to hold the case parts on the CNC mill.
	Display has been chosen as a 16 row 2 line LCD dot matrix used on many other products. Backlight will be controlled by PWM.
	Buttons (x4) to access the menus will be capacitive touch switches mounted just under the case lid. These are highly reliable and as they do not need cutouts do not allow the entry of dust.
	A RS232 “customer” serial port on a DB9F connector.
	The Arduino Mega already has a serial USB port that can be used for uploading firmware.

	Comms
	Has two RJ45 sockets for flexibility in daisy chaining.
	Has links to enable RS485 line level biasing. Probably will not be used.
	Has link for RS485 termination. Likely permanently installed.

	Control Unit Software
	The display shall:
	Show a idle message when it is not executing a command.
	Show the currently executing command.
	Display error messages until they are explicitly cleared with a button press.
	Configuration will be a number of registers containing either numbers or bitwise flags
	These will be stored in EEPROM with appropriate corruption checks.
	Accessed by writing a value into a register by index.
	There is no error checking on writing a particular value. Bad values are simply ignored.
	The list of configuration items will comprise some for development only and a few for the installer. The user will not be able to access them.
	The set of items will depend on the firmware version installed, they are not very friendly for access outside the firmware.
	There will be no menu system with buttons to edit the configuration. Another time saving measure.
	Access will be by the “customer” serial port using a protocol defined TBA and via simple commands on the USB serial port.

	Comms
	The Bed Interface and Sensors will be queried as often as practical.
	A comms failure is not immediately flagged as an error. Only if several errors occur in turn is an error flagged. In this state the system will not operate. The error will be cleared when the unit starts communicating again.
	Logging
	For development speed we have decided to not have a timestamp for logging (which requires a realtime clock), but to use the software tick counter, which counts uptime and repeats every 50 days or so.
	The log memory is to be to a segment of the internal EEPROM memory of the device.
	Log entries will be validated with a CRC or similar.
	This memory is specifically not checked for corruption, as repeated erase on the same memory will eventually wear it out and lead to read errors.
	The size will be whatever size of the EEPROM memory is available.
	Commands
	Commands are a small code identifying an action to take in the system. They comprise:
	All stop.
	Drive each motor in either direction for a certain duration. There may be several duration values used.
	Record position from a number of Sensor units as a position preset.
	Drive motors to return to a preset position.
	Set upper/lower limits for those axes with Sensors installed as limits that the MBC will not drive the motor beyond.
	Only one command can run at a time, they only buffered to a single level to avoid the situation where the bed motors keep moving as commands are removed from a buffer and actioned one by one.
	A command can fail due to a variety of causes, including
	No comms with Bed Interface unit or Sensors
	Sensor unit flags error state.
	Sensor data does not indicate movement when it should be moving.
	Sensor data unstable.
	Attempt to move axis beyond soft limits set by Sensor
	Preset position commanded has not been set yet.
	Timeout getting to preset position.
	etc.
	In case of failure, the command just terminates and cuts the relays. A warning message is displayed on the LCD and may also be logged.
	Commands may originate from:
	the “customer” serial port, the protocol will be determined by the customer to suit their automation system, probably binary with a checksum.
	The USB serial port using a simple protocol that we control. Probably `<nnn> cmd` where <nnn> is a decimal number will do.
	For future versions they could also come from the IR receiver or a speech recognition system.

	Bed Interface Unit – Hardware
	Power
	DC 12-24V input on 2.5mm power inlet, centre positive, to suit the stock 12V adaptor.
	Inlet can be replaced at build time by terminal block for 24V DC.
	Diode to give reverse protection.
	Voltage regulator can accept up to 28V DC.
	PTC thermistor for DC feed to other units, will heat up and go high resistance to limit fault current if there is a short on the bus.
	Comms
	Has two RJ45 sockets for flexibility in daisy chaining.
	Has links to enable RS485 line level biasing. Very likely that these will be always in use.
	Has link for RS485 termination. However we anticipate that this unit will be in the middle of the daisy chain so will not need termination.
	Daughter Board
	Bed interface will be via daughter board with all power rails, SPI, serial and spare IO lines to the processor.
	The default daughter board will have an 8 way relay driver and 8 relays with changeover contacts to pads.
	The pads may be connected by simple wire links to either:
	An 8 way DIN socket (specifically to suit a certain common bed controller).
	A DB9M, to suit existing installations,
	Other daughter boards can be built to suit other bed controllers.

	Faults
	The unit will signal a fault on the LED if it has not received a query for a period, nominally 1 second. This allows easy remote faultfinding.
	The unit will monitor the DC input volts and signal an error if it is below a good value, nominally 9V. The voltage regulators will operate down to 6V but this higher threshold allows problems to be flagged.
	If an error is signalled all relays will immediately turn off.
	Tilt Sensor Unit - Hardware
	Comms
	Has two RJ45 sockets for flexibility in daisy chaining.
	Has link for RS485 termination.
	In case the need for opening the box to set the termination link is troublesome, an alternative strategy is to supply a RJ45 plug with the termination resistor only installed, then it is just plugged into the spare socket of the furthest Sensor by the installer.

	Tilt Sensor Unit – Software
	The critical component in the unit is the accelerometer, the ADXL345 has been chosen as a low noise 3 axis device. Evaluation has been done with a known good unit from Sparkfun and several no-name brands from Aliexpress and they have been found to have similar noise & sensitivity behaviour. As we only need to compute tilt (search for “pitch” on the web as they are often used as part of autopilots) they work very well as the calculation is basically the arctan of two components of acceleration. However care is needed to get over the ±90° range required by keeping track of the sign of the components.
	I have verified that the typical motors on beds (screw driven linear actuators) do not add significant noise to the accelerometer data that cannot be removed with averaging and filtering.
	Also I have set up a simple simulated bed with a motor and accelerometer, and driven it between two preset positions, and the position was always within a degree of the desired. I am confident that his can be improved with a better mechanical setup of a bed and perhaps better filtering.

	We intend to test the system with the same model of bed that the customer will be using.

	To check for correct operation of the accelerometer, several options are possible:
	Verify that the device is sending data at the correct rate. This catches the fault where the device has simply stopped working.
	Verify the acceleration components are reasonable, in particular the scalar magnitude should be 1g, assuming that it is actually accelerating, which would be unusual on a bed. The only problem might be vibration, which can be processed as acceleration, but I could not see any noise due to vibration when the motors were moving, either the filtering removes it or the motors are inherently mechanically quiet.