Scanning and Infra Red Remote Motorised Bed Controller 2023

IR Remote Control
The IR Remote control function uses the RC5 protocol, which has a 5 bit (0..31) System Code and a 8 bit (0..255) Command Code. The system codes are chosen so as to be unlikely to be used by other entertainment equipment.

Serial Port Control
This uses the same System & Command Codes as the IR remote system. Data sent at 9600/8N1.

Codes sent by serial data are sent as a 3 byte sequence:
Byte 1 – System code
Byte 2 – Command code
Byte 3 – Checksum consisting of 2's complement of the sum of the previous 2 bytes. A simple check on reception is then to sum all 3 bytes, only acting on the command if the sum (ignoring carries) is zero.

So for example the command “Leg Up” is sent as 0x30 0x03 0xcd. The sum of these 3 bytes is 0x100, so 0x00 in 8 bits.

Commands are not acknowledged at present but this can be done if required, with a status code, etc.

Safety Measures
1. As a precaution against mis-operation, the controller must be “awake” to act on any command. The wakeup command wakes the controller for 10 seconds before it sleeps again.
2. As a further precaution, the commands to learn a position must be repeated 3 times within 3 seconds. This is addition to the controller needing to be awake.
3. Commands are not queued if the controller is busy, they are simply discarded, unless the command is a repeat of the previous motion command. This allows a stream of such commands to move the bed without stopping and starting.
4. The Stop command will always be acted on immediately.

SBC2023 – RC5 Codes

Motion	System Code	Command Code
Continuous motion while IR is received plus preset run-on (note 1)
		Hexadecimal		Decimal
Head/Backrest up	30 (decimal)	0x01		1
Head/Backrest Down		0x02		2
Leg/Foot up		0x03		3
Leg/Foot down		0x04		4
Bed Up		0x05		5
Bed Down		0x06		6
Tilt/Trendelenburg up		0x07		7
Tilt/ Trendelenburg down		0x08		8
Safety
Stop		0x09		9
Wakeup		0x00		0
Go to Preset positions
Preset 1	30	0x11		17
Preset 2		0x12		18
Preset 3		0x13		19
Preset 4		0x14		20

Programming instructions (note 2)
Learn Preset 1		33 Not sure how you can have a system code of 33 since this is only 5 bits, so 0..31.	0x11	17
Learn Preset 2			0x12	18
Learn Preset 3			0x13	19
Learn Preset 4			0x14	20

Notes
1. Use with systems that just give a single command burst such as speech recognition. The motor will run for a preset time each time the command is received. A series of repeated commands might be needed to reach a desired position.
2. Programming instructions must be received 3 times within 2 seconds to be valid and responded to.